SECTION 11 73 00

HOSPITAL BEDS AND STRETCHERS

Display hidden notes to specifier. (Don't know how? [Click Here](https://www.arcat.com/sd/display_hidden_notes.shtml))

*Copyright 2016 - 2024 ARCAT, Inc. - All rights reserved*

\*\* NOTE TO SPECIFIER \*\* Hillrom, a Baxter Company; Hospital beds and stretchers; health care industry products and solutions.  
This section is based on the products of Hillrom, a Baxter Company, which is located at:1069 State Route 46 E.Batesville, IN 47006Toll Free Tel: 812-934-7777Fax: 812-934-8189Email: [request info (Aman.Agrawal@hillrom.com)](https://arcat.com/rfi?action=email&company=Hillrom%252C%252Ba%252BBaxter%252BCompany&message=RE%253A%2520Spec%2520Question%2520(11731hil)%253A%2520&coid=49149&spec=11731hil&rep=&fax=812-934-8189)  
Web: <https://www.hillrom.com> | <https://www.hillrom.com/construction>   
  
 [ [Click Here](https://arcat.com/company/hillrom-a-baxter-company-49149) ] for additional information.  
  
Hillrom is a leading worldwide manufacturer and provider of medical technologies and related services for the health care industry, including patient support systems, safe mobility and handling solutions, non-invasive therapeutic products for a variety of acute and chronic medical conditions, medical equipment rentals, surgical products and information technology solutions. Hillrom's comprehensive product and service offerings are used by health care providers across the health care continuum and around the world in hospitals, extended care facilities and home care settings to help enhance the safety and quality of patient care.  
Hillrom: Enhancing outcomes for patients and their caregivers.

1. GENERAL
   1. SECTION INCLUDES

\*\* NOTE TO SPECIFIER \*\* Delete items below not required for the project.

* + 1. Hospital birthing beds.
    2. Hospital bariatric beds.
    3. Hospital ICU beds.
    4. Hospital medical surgical beds.
    5. Extended care beds.
    6. Stretchers.
  1. RELATED SECTIONS
     1. Section 27 52 13 - Patient Monitoring Equipment.
     2. Section 11 73 00 - Patient Care Equipment.
     3. Section 11 72 13 - Patient Examination Equipment.
     4. Section 11 76 00 - Operating Room Equipment and Management Systems.
     5. Section 27 52 00 - Health Care Communications and Monitoring Systems.
  2. REFERENCES

\*\* NOTE TO SPECIFIER \*\* Delete references from the list below that are not actually required by the text of the edited section.

* + 1. American National Standards Institute (ANSI):
       1. ANSI/AAMI ES60601-1 - Medical Electrical Equipment-Part 1: General requirements for basic safety and essential performance.
       2. ANSI/UL 1998 - Standard for Software in Programmable Components.
    2. Canadian Standards Association:
       1. CSA C222.2 No. 205 - Signal Equipment.
       2. CSA C22.2 No. 601.1 - Medical Electrical Equipment Part 1: General Requirements for Safety.
       3. CSA C22.2 No. 60601-1. - Medical Electrical Equipment Part 1: General Requirements for Safety. Third edition.
       4. CSA/C22.2 no. 60601-2-38 - Medical Electrical Equipment - Part 2: Particular Requirements for the Safety of Electrically Operated Hospital Beds.
    3. International Committee on Radio Frequency Interference (CISPR):
       1. CISPR 11 - Industrial, Scientific, and Medical Equipment - Radio-Frequency Disturbance Characteristics - Limits and Methods of Measurement.
    4. International Organization for Standardization (ISO):
       1. ISO 10993-1 - Biological evaluation of medical devices -- Part 1: Evaluation and testing within a risk management process.
       2. ISO 10993-5 - Biological evaluation of medical devices -- Part 5: Tests for in vitro cytotoxicity.
       3. ISO 10993-10 - Biological evaluation of medical devices -- Part 10: Tests for irritation and skin sensitization.
       4. ISO 14971 - Medical devices -- Application of risk management to medical devices.
    5. Ministry of Economy, Trade and Industry in Japan:
       1. DENAN PSE Law - Product Safety Electrical Appliance and Material. Compliance with electrical appliance safety for Japan.
       2. J55001 (H20)-Standard for Noise Intensity.
    6. European Committee for Standardization (EN)
       1. EN ISO 9001 - Quality management systems. Requirements.
       2. EN ISO 13485 - Medical devices. Quality management systems. Requirements for regulatory purposes.
       3. EN-60601-1 - Medical electrical equipment. General requirements for basic safety and essential performance.
       4. EN 60601-2-52 - Medical electrical equipment. Particular requirements for the basic safety and essential performance of medical beds.
    7. European Union, OSHA Registration, Evaluation, Authorization and Restriction of Chemicals (REACH):
       1. REACH Directive 1907/2006
    8. European Union, Restriction of Hazardous Substances Directive (RoHS):
       1. RoHS Directive 2002/95/EC
    9. European Union: Waste Electrical and Electronic Equipment (WEEE):
       1. WEEE Directive 2002/96/EC.
    10. International Electrotechnical Commission (IEC)
        1. IEC 60601-1 - Medical electrical equipment - Part 1: General requirements for basic safety and essential performance.
        2. IEC 60601-1-2 - Medical electrical equipment - Part 1-2: General requirements for basic safety and essential performance - Collateral Standard: Electromagnetic disturbances - Requirements and tests.
        3. IEC 60601-1-4 - Medical electrical equipment - Part 1-4: General requirements for safety - Collateral Standard: Programmable electrical medical systems.
        4. IEC 60601-1-6 - Medical electrical equipment - Part 1-6: General requirements for basic safety and essential performance - Collateral standard: Usability.
        5. IEC-60601-2-38 - Medical Electrical Equipment - Part 2-38: Particular Requirements for the Safety of Electrically Operated Hospital Beds.
        6. IEC 60601-2-46 - Medical electrical equipment - Part 2-46: Particular requirements for the basic safety and essential performance of operating tables.
        7. IEC-60601-2-52 - Medical electrical equipment - Part 2-52: Particular requirements for the basic safety and essential performance of medical beds.
        8. IEC 61000-3-2 - Electromagnetic compatibility. Limits - Limits for harmonic current emissions (equipment input current leas than or equal to 16 A per phase).
        9. IEC 61000-3-3 - Electromagnetic compatibility. Limits - Limitation of voltage changes, voltage fluctuations and flicker in public low-voltage supply systems, for equipment with rated current less than or equal to 16 A per phase and not subject to conditional connection.
        10. IEC 61000-4-2 - Electromagnetic compatibility. Part 4-2: Testing and measurement techniques - Electrostatic discharge immunity test.
        11. IEC 61000-4-3 - Electromagnetic compatibility. Part 4-3: Testing and measurement techniques - Radiated, radio-frequency, electromagnetic field immunity test.
        12. IEC 61000-4-4 - Electromagnetic compatibility. Part 4-4: Testing and measurement techniques - Electrical fast transient/burst immunity test.
        13. IEC 61000-4-5 - Electromagnetic compatibility. Part 4-5: Testing and measurement techniques - Surge immunity test.
        14. IEC 61000-4-6 - Electromagnetic compatibility. Part 4-6: Testing and measurement techniques - Immunity to conducted disturbances, induced by radio-frequency fields.
        15. IEC 61000-4-8 - Electromagnetic compatibility. Part 4-8: Testing and measurement techniques - Power frequency magnetic field immunity test.
        16. IEC 61000-4-11 - Electromagnetic compatibility. Part 4-11: Testing and measurement techniques - Voltage dips, short interruptions and voltage variations immunity tests.
    11. Underwriters Laboratory (UL):
        1. UL 1069 - UL Standard for Safety Hospital Signaling and Nurse Call Equipment.
        2. UL 60601-1 - Medical Electrical Equipment, Part 1: General Requirements for Safety.
  1. SUBMITTALS
     1. Submit under provisions of Section 01 30 00 - Administrative Requirements.
     2. Product Data: Manufacturer's data sheets on each product to be used, including:
        1. Source quality certificates.
        2. Preparation instructions and recommendations.
        3. Storage and handling requirements and recommendations.
        4. Installation methods.
        5. Maintenance and operations data.
     3. Shop Drawings: Include system components, utility requirements and connections, relationship with adjacent construction. Include required clearances and access for servicing.
        1. Communications wire labeling schedules.
        2. Communications wiring diagrams: Access points.
  2. QUALITY ASSURANCE
     1. Standards Compliance:
        1. EN ISO 9001 - Quality management systems. Requirements.
        2. EN ISO 13485 - Medical devices. Quality management systems. Requirements for regulatory purposes.
     2. Regulatory Requirements: Comply with requirements of authorities having jurisdiction and applicable codes at the location of the project.
     3. Manufacturer Qualifications: Minimum 5 years experience manufacturing similar products.
     4. Installer Qualifications: Minimum 2 years experience installing similar products.
  3. DELIVERY, STORAGE, AND HANDLING
     1. Deliver and store products in manufacturer's unopened packaging bearing the brand name and manufacturer's identification until ready for installation.
     2. Comply with manufacturer's recommendations. Handle materials to avoid damage.
  4. PROJECT CONDITIONS
     1. Maintain environmental conditions (temperature, humidity, and ventilation) within limits recommended by manufacturer for optimum results. Do not install products under environmental conditions outside manufacturer's recommended limits.
  5. WARRANTY
     1. Provide manufacturer's standard limited warranty.

1. PRODUCTS
   1. MANUFACTURERS
      1. Acceptable Manufacturer: Hillrom, a Baxter Company, which is located at:1069 State Route 46 E.Batesville, IN 47006Toll Free Tel: 812-934-7777Fax: 812-934-8189Email: [request info (Aman.Agrawal@hillrom.com)](https://arcat.com/rfi?action=email&company=Hillrom%252C%252Ba%252BBaxter%252BCompany&message=RE%253A%2520Spec%2520Question%2520(11731hil)%253A%2520&coid=49149&spec=11731hil&rep=&fax=812-934-8189);Web: <https://www.hillrom.com> | <https://www.hillrom.com/construction>

\*\* NOTE TO SPECIFIER \*\* Delete one of the following two paragraphs; coordinate with requirements of Division 1 section on product options and substitutions.

* + 1. Substitutions: Not permitted.
    2. Requests for substitutions will be considered in accordance with provisions of Section 01 60 00 - Product Requirements.

\*\* NOTE TO SPECIFIER \*\* Delete article if not required.

* 1. HOSPITAL BIRTHING BEDS
     1. Basis of Design: Hillrom Smart Beds and Surfaces; Affinity 4 Birthing Beds
        1. Standards Compliance:
           1. ANSI/AAMI ES60601-1.
           2. CSA C22.2 No. 205.
           3. CSA C22.2 No. 60601-1.
           4. DENAN Law Japan.
           5. EN 60601-1.
           6. IEC 60601-1-2.
           7. IEC 60601-2-38.
           8. IEC 60601-2-52.
           9. ISO 10993-1.
           10. J55001 (H20) Japan.

\*\* NOTE TO SPECIFIER \*\* Delete models not required.

* + - 1. Model AF400:
         1. Prevention (foam) surface.
         2. Lift-off foot section.
      2. Model AF450:
         1. Prevention (foam) surface.
         2. Nurse call, lighting, entertainment and UTV.
         3. Lift-off foot section.
      3. Model AF500:
         1. Air package (lumbar and seat).
         2. Lift-off foot section.
      4. Model AF550:
         1. Air package (lumbar and seat).
         2. Nurse call, lighting, entertainment and UTV.
         3. Lift-off foot section.
      5. Model AF600:
         1. Prevention (foam) surface.
         2. Stow and Go foot section.
      6. Model AF650:
         1. Prevention (foam) surface.
         2. Nurse call, lighting, entertainment and UTV.
         3. Stow and Go foot section.
      7. Model AF700:
         1. Air package (lumbar and seat).
         2. Stow and Go foot section.
      8. Model AF750:
         1. Air package (lumbar and seat).
         2. Nurse call, lighting, entertainment and UTV.
         3. Stow and Go foot section.

\*\* NOTE TO SPECIFIER \*\* Delete seat cut not required.

* + - 1. Seat Cut: Straight.
      2. Seat Cut: V.

\*\* NOTE TO SPECIFIER \*\* Delete headboards not required.

* + - 1. Headboard: Conaorary (wood finish).
      2. Headboard: Freedom Hill (wood finish).
      3. Headboard: Liberty Hill (wood finish).
      4. Standard Features:
         1. V-cut or straight cut surface.
         2. Removable foot section.
         3. Battery back-up.
         4. Powered foot section.
         5. Casters: Four wheel dual locking.
         6. OneStep siderail release.
         7. Plastic head panel.
         8. Automatic nightlight.
         9. Lock-out controls.
         10. Manual trend-like positioning.
         11. Instant CPR.
         12. Instant labor grips.
         13. Automatic seat tilt (15 degrees).
         14. Siderail controls.
         15. IV Pole.
         16. Hinged foot section mattress.
         17. Central brake and steering.
         18. Fluid basin.
         19. Mattress:

Length: 78 in (1981 mm).

Width: 35 in (889 mm).

Thickness (head section): 5 in (127 mm).

Thickness (foot section): 3.5 in (89 mm).

* + - * 1. Safe Working Load: 480 lbs (227 kg).

\*\* NOTE TO SPECIFIER \*\* Delete bed weight not required.

* + - * 1. Bed Weight: 480 lbs (217 kg) lift-off foot section.
        2. Bed Weight: 511 lbs (232 kg) Stow and Go foot section.
      1. Optional Features:

\*\* NOTE TO SPECIFIER \*\* Delete optional features not required.

* + - * 1. Calf Supports.
        2. Push handles with IV hook.
      1. Finish Options:

\*\* NOTE TO SPECIFIER \*\* Delete finish options not required.

* + - * 1. Natural Maple.
        2. Wild Cherry.
        3. Dark Cherry.
        4. Colonial Cherry.
        5. Montana Walnut.
        6. Honey Maple.
        7. Light Oak.
        8. Medium Oak.
      1. Wood Grain Laminate Options:

\*\* NOTE TO SPECIFIER \*\* Delete laminate options not required.

* + - * 1. Natural Maple.
        2. Wild Cherry.
        3. Dark Cherry.
        4. Shaker Cherry (matches Colonial Cherry in wood finish).
        5. Montana Walnut.
        6. Honey Maple.
        7. Light Oak.
        8. Medium Oak.
      1. Color Laminate Options:

\*\* NOTE TO SPECIFIER \*\* Delete laminate options not required.

* + - * 1. Almond Leather.
        2. Natural Legacy.
        3. Antique White.
        4. White Nebula.
        5. Light Neutral.
      1. Overall Frame Dimensions:
         1. Frame Width:

Side Rails Up: 39 in (991 mm).

Side Rails Down: 36 in (914 mm).

* + - * 1. Frame Length: 92 inches (2337 mm).
      1. Bed Height:
         1. Low (with mattress): 22 in (560 mm).
         2. High (with mattress): 38 in (970 mm).
         3. Seat Section in Trend-Like Position: 40 in (1016 mm) maximum.
      2. Bed Clearance Under Frame: 5 in (127 mm).
      3. Bed Articulations:
         1. Head Elevation: 63 degrees.
         2. Pelvic Tilt: Automatic 5 degrees.
         3. Trend-Like Positioning: 8 degrees.
      4. Caster: 6 in (152 mm) diameter.

\*\* NOTE TO SPECIFIER \*\* Delete article if not required.

* 1. HOSPITAL BARIATRIC BEDS

\*\* NOTE TO SPECIFIER \*\* Delete paragraph if not required.

* + 1. Basis of Design: Hillrom Smart Beds and Surfaces; Compella Bariatric Beds.
       1. Standards Compliance:
          1. CAN/CSA C22.2 No. 60601-1:08
          2. IEC 60601-1-2.
          3. IEC/EN/ANSI/AAMI ES60601-1
          4. IEC/EN 60601-1
          5. IEC/EN 60601-1-2
          6. IEC/EN 60601-1-6
          7. IEC/EN 60601-2-52
          8. EN ISO 9001
          9. EN 13485
          10. REACH Directive 1907/2006
          11. RoHS Directive 2002/95/EC
          12. WEEE Directive 2002/96/EC
       2. Standard Features:
          1. Powered Width Expansion: 40 to 50 in (1016 to 1270 mm).
          2. Powered Length Adjustment: 80 to 88 in (2032 to 2235 mm).
          3. Single mode bed exit.
          4. Digital head of bed indicator and alert.
          5. Integrated scale.
          6. Battery backup.
          7. Night light.
          8. Integrated caregiver and patient controls.
          9. Slideguard technology.
          10. Patient Weight Range: 250 to 1000 lbs (113 to 454 kg).
          11. Safe Working Load: 1100 lbs (498.9 kg).

\*\* NOTE TO SPECIFIER \*\* Delete bed weight not required.

* + - * 1. Bed Weight: 838 lbs (380 kg), without IntelliDrive XL transport system.
        2. Bed Weight: 948 lbs (430 kg), with IntelliDrive XL transport system.
      1. Surface Options:

\*\* NOTE TO SPECIFIER \*\* Delete surfaces not required.

* + - * 1. Surface: Foam with air supply unit.
        2. Surface: Low air loss surface with turn assist and air supply unit.

\*\* NOTE TO SPECIFIER \*\* Delete surface add on if not required.

* + - 1. Surface Add On: Continuous lateral rotation therapy.
      2. Optional Features:

\*\* NOTE TO SPECIFIER \*\* Delete options not required.

* + - * 1. Three mode bed exit alert with alert silence (Patient position monitoring).
        2. Intellidrive XL powered transport.
        3. Patient Helper: Full assembly.
        4. Pendant.
        5. Line manager.
        6. IV Pole.
        7. Enabled smart bed connectivity.
        8. SNC - Siderail nurse call. Requires enabled smart bed connectivity.
        9. Pressure transducer holder.
      1. Overall Frame Dimensions:
         1. Frame Width (extenders retracted): 43 in (1092 mm).
         2. Frame Length:

Normal 91 in (2311 mm).

Extended 98 in (2489 mm).

* + - 1. Sleep Deck Dimensions:
         1. Width 40 to 50 in (1016 to 1270 mm).
         2. Length: 80 to 88 in (2032 to 2235 mm).
      2. Bed Height:
         1. Low Position Floor to Sleep Deck: 18-1/2 in (470 mm).
         2. High Position Floor to Sleep Deck: 26-3/4 in (680 mm).
         3. Transport Position: 17 in (432 mm).
      3. Surface Height Inflated:
         1. Foam: 8 in (203 mm).
         2. Low Air Loss: 8.5 in (215 mm).
      4. Scale System:
         1. Accuracy: Plus or minus 2.2 lbs (1 kg) or 1 percent of patient weight whichever is greater.
         2. Weigh in any position.
         3. Scale repeatability 0.2 percent of patient weight.
      5. Bed Articulations:
         1. Head Elevation: 0 to 50 degrees.
         2. Knee Elevations: 0 to 30 degrees.
         3. Trendelenburg and Reverse Trendelenburg: 0 to 9 degrees.
      6. Caster Size: 6 in (152 mm).

\*\* NOTE TO SPECIFIER \*\* Delete paragraph if not required

* + 1. Basis of Design: Hillrom Smart Beds and Surfaces; TotalCare Bariatric Plus Hospital Beds.
       1. Standards Compliance:
          1. CSA C22.2 No. 601.1.
          2. IEC 60601-1-2.
          3. IEC 60601-1-4.
          4. IEC/EN 60601-2-38.
          5. IEC/EN 60601-1.

\*\* NOTE TO SPECIFIER \*\* Delete models not required.

* + - 1. Model TCB-765:
         1. Low air loss surfaces.
         2. Advanced microclimate technology.
         3. Rotation and percussion/vibration permanently installed.
         4. Low air loss module permanently installed.

\*\* NOTE TO SPECIFIER \*\* Delete if not required.

* + - * 1. IV Pole.
      1. Model TCB-655:
         1. Low air loss surfaces.
         2. Advanced microclimate technology.
         3. Low air loss module permanently installed.

\*\* NOTE TO SPECIFIER \*\* Delete if not required.

* + - * 1. IV Pole.
      1. Model TCB-155:
         1. Foam surface (3 layer).
         2. Siderail extender.

\*\* NOTE TO SPECIFIER \*\* Delete if not required.

* + - * 1. IV Pole.
      1. Standard Features:
         1. Surface Width: 40 in (1016 mm).
         2. Full chair position.
         3. Chair egress with stand assist.
         4. SlideGuard frame technology.
         5. FlexAfoot retractable foot control.
         6. Head and foot panel.
         7. 30 and 45 degree head of bed alarm.
         8. Smart Bed Ready: Sidecom required.
         9. Caster System: 5 in (127 mm) urethane central locking (not available with IntelliDrive).
         10. Both siderails with point of care touch screen.
         11. Bed exit alarm.
         12. Point-of-care siderail controls.
         13. Color touch screen graphic caregiver interface.
         14. Line of site angle indicators.
         15. On-board help.
         16. On-board graphic histories.
         17. Radiolucent sleep deck.
         18. Line managers.
         19. Wallguard roller bumper system.
         20. Drainage bag holders.
         21. Hydraulic foot pump frame articulation.
         22. Hands free emergency CPR and Trendelenburg.
         23. TuckAway siderails with OneStep siderail release.
         24. Night light.
         25. Electric Trendelenburg.
         26. In-bed scale.
         27. Boost function.
         28. Head board.
         29. Foot board.
      2. Optional Features:

\*\* NOTE TO SPECIFIER \*\* Delete optional features not required.

* + - * 1. Accessory outlet.
        2. NUL - Siderail nurse call, universal TV and lighting controls (Does not include bed status).
        3. NSC- Bed controls only. Bed exit alarm to nurse call system; there are no nurse call buttons in the rails.
        4. Intellidrive powered transport mechanism.
        5. Caster System: 6 in (152 mm) Urethane Central Locking (Must be used with Intellidrive).
        6. Twin Caster: 5 in (127 mm): Not available with Intellidrive.
        7. Nylon coverlet without silver (launderable).
        8. IV Pole.
        9. Patient Helper: Complete Assembly.
        10. Patient Helper: Base Only.
        11. Navicare System Ready Local Alarm On and Off: Requires Scale or PPM. For Navicare System applications only.
      1. Overall Frame Dimensions:
         1. Width (side rails up): 44.5 in (1130 mm) maximum.
         2. Length: 81-1/2 in (2070 mm) minimum.
      2. Sleep Deck Dimensions:
         1. Surface Width: 40 in (1016 mm).
         2. Surface Length: 72 to 84 in (1829 to 2134 mm).
         3. Surface Height:

Foam: 7 in (178 mm).

Low Air Loss: 11 in (279 mm).

* + - * 1. Sleep Deck to Floor Low Position: 19 in (483 mm).
        2. Sleep Deck to Floor High Position: 38 in (965 mm).
      1. Patient Weight Capacity: 500 lbs (227 kg).

\*\* NOTE TO SPECIFIER \*\* Delete article if not required.

* 1. HOSPITAL ICU BEDS
     1. Basis of Design: Hillrom Smart Beds and Surfaces; Progressa Bed Systems.
        1. Standards Compliance:
           1. CAN/CSA-22.2 NO. 60601-1.
           2. IEC 60601-1-2.
           3. ES60601-1.
           4. IEC/EN60601-1.
           5. IEC/EN60601-2-52.
           6. IEC/EN60529

Cleanability and Infection Control: IPX4.

* + - 1. Physical Attributes:
         1. Overall Frame Dimensions:

Width: 40-1/2 in (1029 mm).

Length: 88 in (2235 mm).

Top of Seat Section to Floor:

Highest: 35 in (889 mm).

Lowest: 16-1/2 in (419 mm).

* + - * 1. C-Arm Clearance: 7 to 15 in (178 to 381 mm).
        2. Bed Articulations

Head: 67 degrees.

Thigh: 30 degrees.

Foot: 70 degrees.

Trendelenburg: 13 degrees.

Tilt Table Function (Reverse Trendelenburg): -18 degrees.

* + - * 1. Voltages: 100, 110, 115, 120, 127, 220, 230, and 240.
        2. Overall Weight: 800 lbs (363 kg).
      1. Standard Features:
         1. SlideGuard frame technology.
         2. Patient controls backlighting.
         3. FlexAfoot retractable foot control.
         4. Advanced MicroClimate Technology (permanently installed).
         5. 30 to 45 degree head of bed alarm.
         6. Bed controls.
         7. In-bed scale.
         8. Line managers.
         9. Three mode bed exit alarm (PPM).
         10. Caregiver pendant.
         11. Standard head and foot panel.
         12. Foot controls module.
         13. Night light.
         14. Oxygen tank holder module.
         15. Steering caster.
         16. Accessory outlet.
         17. Hands free emergency CPR, Trendelenburg.
         18. NaviCare Ready: Local alarm On/Off.
         19. Point-of-Care siderails control.
         20. Both siderails with point-of-care touch screen.
         21. Smart Bed Ready (Sidecom required).
         22. Caster System: 6 in (152 mm) urethane central locking.
         23. Obstacle detection system on both sides of the bed.
         24. Maximum Patient Weight: 500 lbs (226 kg).
         25. Safe Working Load: 650 lbs (295 kg).
         26. Drainage bag holders.
         27. Four IV sockets.

\*\* NOTE TO SPECIFIER \*\* Delete dining chair if not required.

* + - * 1. Dining chair.

\*\* NOTE TO SPECIFIER \*\* Delete the surface option not required.

* + - * 1. Prevention Surface: Non-powered air.
        2. Therapy Surface: Powered air.

\*\* NOTE TO SPECIFIER \*\* Rotation requires the therapy surface. Delete if not required.

CLRT: Rotation.

\*\* NOTE TO SPECIFIER \*\* Percussion/vibration requires rotation. Delete if not required.

Percussion/Vibration.

\*\* NOTE TO SPECIFIER \*\* Delete the two nurse call system options not required.

* + - * 1. SNC: Siderail nurse call system.
        2. NUL: Siderail nurse call system, universal TV and lighting controls.
        3. NSC: Bed exit alarm nurse call system. No siderail nurse call system.
      1. Exclusive Hill-Rom Features:
         1. Corner steer.
         2. SlideGuard frame technology.
         3. Hands free foot controls module.
         4. Line managers.
         5. Caregiver pendant.
         6. Integrated O2 tank holder.
         7. Audible "brake not set" alarm.
         8. Scale and PPM. 3 mode bed exit auto reset w/ alarm silence.
         9. Digital head of bed display and alarm.
         10. Battery back up. On all bed articulations.
         11. Flex-A-foot Integrated bed extender. Foot ended.
         12. No controls in footboard.
         13. Accessory outlet.
         14. Built-in quick reference guide.
      2. Exclusive Hill-Rom Optional Features:
         1. Full chair egress.

\*\* NOTE TO SPECIFIER \*\* Requires full chair egress. Delete if not required.

StayInPlace patient migration management system.

* + - * 1. WatchCare incontinence detection system.

\*\* NOTE TO SPECIFIER \*\* Delete options not required or delete entire paragraph if no options are required.

* + - 1. Options:
         1. Fifth wheel steering.
         2. IntelliDrive powered transport mechanism.
         3. Transport shelf.
         4. Wireless.

\*\* NOTE TO SPECIFIER \*\* Delete one of the following two subparagraphs or delete both if a permanent IV pole is not required.

* + - * 1. Permanent IV pole, right side.
        2. Permanent IV pole, left side.

\*\* NOTE TO SPECIFIER \*\* Delete article if not required.

* 1. HOSPITAL MEDICAL SURGICAL BEDS

\*\* NOTE TO SPECIFIER \*\* Delete paragraph if not required.

* + 1. Basis of Design: Hillrom Smart Beds and Surfaces; Centrella� Smart+ medical-surgical beds.
       1. Standards Compliance:
          1. CAN/CSA C22.2 No. 60601-1.
          2. ANSI/AAMI ES60601-1
          3. IEC 60601-1.
          4. IEC 60601-1-2.
          5. IEC 60601-1-6.
          6. IEC 60601-2-52.
          7. IEC 62304.

\*\* NOTE TO SPECIFIER \*\* Delete: models not required.

* + - 1. Model Centrella base frame:

\*\* NOTE TO SPECIFIER \*\* Delete optional items not required or

* + - * 1. Optional Features:

Optimized Patient Safety Features:

Centrella max Powered Air Surface: Includes blower, comfort controls, turn assist and focused Microclimate management.

Centrella pro Surface.

Centrella core Surface.

Enhanced Patient Satisfaction Features:

Comfort Edition: 40 in (1016 mm) width.

Patient pendant.

Second patient pendant.

Nurse call.

Universal TV and lighting controls.

Advanced Caregiver-Focused Technology Features:

5th wheel steering.

Second graphical caregiver interface (GCI) touchscreen.

Accessory outlet.

Additional line management.

Patient Safety Applications:

NaviCare patient safety.

EarlySense contact-free continuous monitoring.

WatchCare incontinence detection system.

* + - 1. Model Base Frame + PAK255:
         1. Additional Standard Features:

Verbal safety prompts.

USB charging.

Patient pendant.

* + - * 1. Optional Features:

Optimized Patient Safety Features:

Centrella max Powered Air Surface: Includes blower, comfort controls, turn assist and focused Microclimate management.

Centrella pro Surface.

Centrella core Surface.

IllumiGuide Siderail Handgrip: Available with SafeView+ System.

Enhanced Patient Satisfaction Features:

Comfort Edition 40 in (1016 mm) width.

Second patient pendant.

Nurse call.

Universal TV and lighting controls.

Advanced Caregiver-Focused Technology Features:

SafeView+ System: Available with IllumiGuide siderail handgrip.

5th wheel steering.

IntelliDrive powered transport.

Second graphical caregiver interface (GCI) touchscreen.

Accessory outlet.

Additional line management.

Patient Safety Applications:

NaviCare patient safety.

EarlySense contact-free continuous monitoring.

WatchCare incontinence detection system.

* + - 1. Model Centrella Base Frame + PAK355:
         1. Additional Standard Features:

Centrella max powered air surface.

IllumiGuide siderail handgrip.

Safeview+ System.

Second graphical caregiver interface(GCI) touchscreen.

Verbal safety prompts.

USB charging.

Patient pendant.

* + - * 1. Optional Features:

Enhanced Patient Satisfaction Features:

Comfort Edition 40 in (1016 mm) width.

Second patient pendant.

Nurse call.

Universal TV and lighting controls.

Advanced Caregiver-Focused Technology Features:

5th wheel steering.

IntelliDrive powered transport.

Accessory outlet.

Additional line management.

Patient Safety Applications:

NaviCare patient safety.

EarlySense contact-free continuous monitoring.

WatchCare incontinence detection system.

* + - 1. Standard Features:
         1. Optimized Patient Safety Features:

Enhanced 3-mode bed exit.

Motion-activated nightlight.

One-button patient stand assist.

Advanced SlideGuard technology.

Foley position limit.

Patient HOB angle display.

Trendelenburg and reverse Trendelenburg.

Dampen emergency CPR.

* + - * 1. Enhanced Patient Satisfaction Features:

Device storage area.

Integrated urinal holders.

FlexAfoot powered bed extension.

* + - * 1. Advanced Caregiver-Focused Technology Features:

Graphical caregiver interface (GCI) touchscreen (right side).

Bed exit silence.

Brake-not-set alert.

One button chair.

Scale (kg/lbs).

Battery backup.

Digital angle indicator.

Four corner brake and steer.

Line and cord management.

Four IV pole sockets (located at each corner).

Integrated restraint holders.

Lockout controls.

Head-end traction.

Drainage bag holder.

Smart bed ready; SideCom required.

WiFi.

* + - 1. Accessories:
         1. Experience POD device.
         2. Transport shelf.
         3. Patient helper and traction parts.
         4. ISS Transfer pole adaptor.
         5. Standard removable IV pole.
         6. Permanent IV pole.
         7. O2 tank holder.
         8. SideCom breakaway cables.
         9. Ergonomic push handles.
         10. Additional line management.
      2. Overall Frame Dimensions:
         1. Width: 40 in (1016 mm).
         2. Width; Comfort Edition: 43 in (1092 mm).
         3. Length; Foot Section Extended: 98.7 in (2507 mm).
         4. Length; Foot Section Retracted: 86.7 in (2202 mm).
      3. Sleep Deck Dimensions:
         1. Centrella Core:

Thickness: 6 in (152 mm).

Width: 35.5 in (775 mm).

Length Extended: 88 in (2235 mm).

Length Retracted: 76 in (1930 mm).

* + - * 1. Centrella Pro:

Thickness: 7 in (178 mm).

Width: 35.5 or 38.5 in (775 or 978 mm).

Length Extended: 88 in (2235 mm).

Length Retracted: 76 in (1930 mm).

* + - * 1. Centrella Max:

Thickness: 8 in (203 mm).

Width: 35.5 or 38.5 in (775 or 978 mm).

Length Extended: 88 in (2235 mm).

Length Retracted: 76 in (1930 mm).

* + - 1. Bed Height:
         1. Maximum: 32.5 in (825 mm).
         2. Minimum: 14.5 in (368 mm).
      2. Bed Clearance Under Frame: 5.75 in (127 mm).
      3. Safe Working Load: 650 lbs (295 kg).
         1. Patient weight range 70 to 500 lbs (32 to 227 kg).
      4. Bed Articulation:
         1. Head Inclination: 65 degrees.
         2. Thigh Inclination: 25 degrees.
         3. Foot Section Inclination: 50 degrees.
         4. Trendelenburg and Reverse Trendelenburg: 15 degrees.
      5. Caster Size: 5 in (127 mm) with central locking system.

\*\* NOTE TO SPECIFIER \*\* The Envella(r) Air Fluidized Therapy System is a patient management system that combines air fluidized therapy and pressure redistribution technologies together on a low-height frame. The lower body section provides air fluidized therapy, and the upper body section provides pressure redistribution.  
The bed is most applicable in the long term care and acute care, medical, surgical, and critical care settings, but it may be used in other clinical areas as well. Patient Support: Aggressive treatment for multiple or advanced pressure injuries, Treatment for flaps, grafts, burns, or intractable pain, and those at high risk for skin breakdown (to prevent pressure injuries). Delete paragraph if not required.

* + 1. Basis of Design: Hillrom Smart Beds and Surfaces; Envella Air Fluidized Therapy Beds.
       1. When airflow is turned on, room air is drawn into bed base, filtered, heated, and channeled into the bed through a porous diffuser. Air blows upward through the beads putting them in motion, taking on fluid properties on which a patient's body floats. When turned off, beads settle into a mold of body contours creating a solid support system.
       2. Standards Compliance:
          1. IEC 60601-1: Class I equipment.
          2. ANSI/AAMI ES60601-1.
          3. IEC 80601-2-35.
          4. IEC 60601-1-8.
          5. IEC 60601-1-2.
          6. IEC 60601-2-52.

Intensive/critical care (environment 1).

Acute care (environment 2).

Long term care (environment 3).

Outpatient/ambulatory care (environment 5).

* + - * 1. Degree of protection against electric shock: Type B.
        2. Degree of protection against ingress of water: IPX4.
        3. Degree of protection against the presence of flammable anesthetic mixtures: Not for use with flammable anesthetics.
      1. Model P0819A:
      2. Standard Features:
         1. Length with and without Accessories: 92.5 inches (2350 mm) max.
         2. Width, Siderails Removed or Installed: 40 inch (1016 mm) max.
         3. Height with Accessories and Options when Bed is in its Lowest Position: 78 inches (1981 mm).
         4. Distance between inside of headboard and inside of footboard: 83.5 inch (2121 mm).
         5. Air Fluidization Chamber: Silicone coated microsphere (beads).
         6. Width of bead section: 19 inch (483 mm) min.
         7. Length of the bead section: 45 inch (1143 mm) min.
         8. Bed Frame Ground Clearance: 3 inch (76 mm) min.
         9. Height travel Range: 22 to 31.5 inches (559 to 800 mm).
         10. Head Elevation Range: 0 to 60 degrees.
         11. Head Section Panel Height at Lowest Position with Bed Flat: 22 inch (559 mm).
         12. Patient Weight Range: 70 to 350 lb (32 to 159 kg).
         13. Safe Working Load: 450 lb (204 kg).
         14. Total Bed Weight with Safe Working Load: 1874 lb (850 kg).
         15. Sound Pressure Level: Per IEC 60601-2-52.

Normal Fault Free Operation, Excluding Actuator: 65 dBA max.

Short Term/Impulsive Operation: 85 dBA max.

* + - * 1. Sound Pressure Level Per IEC 60601-1-8, Medium Priority Alarm Signal: 74 dBA.
        2. Bead Temperature:

Control Range: 86 to 100 degrees F (30 to 38 degrees C).

Comfort Range: 88 to 94 degrees F (31 to 34 degrees C)

* + - * 1. Patient Scale:

Accuracy and Repeatability: 2.2 lb (0.99 kg) or 1 percent of patient weight, whichever is greater.

* + - 1. Graphical Caregiver Interface (GCI) Control and Panel: Touchscreen at foot of bed. Continuous information about bed features and status. Indicator lights.
         1. View helpful information for the bed functions.
         2. Set Bed Exit and Head Angle Alerts.
         3. Preemptively silence the Bed Exit Alert.
         4. Zero the scale.
         5. Weigh the patient. Calculate BMI. View history.
         6. Adjust the Bed Therapy settings.
         7. Temperature Adjustment: Affects bead section only. 86 to 100 degrees F (30 to 38 degrees C) in one degree F (0.5 degrees C) increments.

Temperature Variations: Up to 3 degrees F (1.7 degrees C).

* + - * 1. Set a Turn Reminder.
        2. Review the history for the different bed functions.
        3. Adjust the GCI's date and time.
        4. Historical data storage on features.
        5. Additional controls: Pendant hand control, footboard, and bed side.
      1. Accessories:
         1. Patient Helper: P0820.
         2. Foot Support Cushion: 195108.
         3. Impervious Sheet: P007012A.
         4. Infusion Support System (ISS): P158A.

Adapter Bracket: P163 for use of ISS.

* + - * 1. IV Pole: P2217A.
        2. Oxygen Tank Holder, Vertical: P276 and P27601.
      1. Environmental Operating Conditions:
         1. Ambient operating temperature range. Variation in temperature specifications could be caused by humidity, restricted air circulation from blankets and pillows, or recent fluid ingress into beads.

Air fluidization Temperature:

Setpoint between 86 and 100 degrees F (30 and 37.8 degrees C): 72 to 78 degrees F (22.2 to 25.6 degrees C).

Setpoint between 86 and 98 degrees F (30 and 36.7 degrees C): 70 to 71 degrees F (21.1 to 21.7 degrees C).

Setpoint between 86 and 96 degrees F (30 and 35.6 degrees C): 68 F to 69 degrees F (20 to 20.6 degrees C).

* + - * 1. Relative Humidity: 20 to 75 percent.
        2. Pressure: 82.8 to 106 kPa.
        3. Altitude: 5500 ft (676 m).

\*\* NOTE TO SPECIFIER \*\* The emissions characteristics of this equipment make it suitable for use in industrial areas and hospitals (CISPR 11 class A). If it is used in a residential environment (for which CISPR 11 class B is normally required) this equipment might not offer adequate protection to radio-frequency communication services. The user might need to take mitigation measures, such as relocating or re-orienting the equipment.

* + - * 1. Electromagnetic Environment: RF energy used only for internal functions. Low emissions will not likely cause interference in nearby electronic equipment.

RF Emissions; CISPR 11: Group 1. Class A.

Harmonic Emissions per IEC 61000-3-2: Not applicable.

Voltage Fluctuations / Flicker Emissions IEC 61000-3-3: Not applicable.

* + - 1. Power Requirements: Rated Voltage: 120 VAC, 60 Hz. Rated Current :10 A.
         1. No user accessible fuses. Envella Air Fluidized Therapy System Service Manual (194345) for fuse ratings and replacement procedures.
      2. Mode of Operation: 2 minutes ON / 18 minutes OFF
      3. Applies Parts in Accordance with IEC60601-1:
         1. Siderails; not including the posts.
         2. Footboard; not including the posts.
         3. Headboard; not including the posts.
         4. Support Surface; Coverlet, Filter sheet, Air wall cover, Foot cushion.
         5. Pendant: Hand control.

\*\* NOTE TO SPECIFIER \*\* Delete article if not required.

* 1. EXTENDED CARE BEDS

\*\* NOTE TO SPECIFIER \*\* Delete paragraph if not required.

* + 1. Basis of Design: Hillrom Smart Beds and Surfaces; Hillrom 100 Low Beds.
       1. Standards compliance:
          1. CAN/CSA C22.2 No. 601.1.
          2. CSA/C22.2 no. 60601-2-38.
          3. IEC 60601-1.
          4. IEC 60601-1-2.
          5. IEC 60601-2-38.
          6. UL 60601-1.

\*\* NOTE TO SPECIFIER \*\* Delete models not required.

* + - 1. Model LB-NONE:

\*\* NOTE TO SPECIFIER \*\* Delete frame color not required.

* + - * 1. Frame Color: Stone.
        2. Frame Color: Silver.
        3. Side Rails: No.

\*\* NOTE TO SPECIFIER \*\* Delete headboard and footboard not required

* + - * 1. Headboard and Footboard: Standard.
        2. Headboard and Footboard: Sleek.

\*\* NOTE TO SPECIFIER \*\* Delete finial style not required.

* + - * 1. Finial Style: Standard.
        2. Finial Style: Traditional.

\*\* NOTE TO SPECIFIER \*\* Delete one of the two following subparagraphs

* + - * 1. Assembled.
        2. Boxed.
      1. Model LB-HALFF:

\*\* NOTE TO SPECIFIER \*\* Delete frame color not required.

* + - * 1. Frame Color: Stone.
        2. Frame Color: Silver.
        3. Assist Rails: 1/2 Fixed.

\*\* NOTE TO SPECIFIER \*\* Delete headboard and footboard not required

* + - * 1. Headboard and Footboard: Standard.
        2. Headboard and Footboard: Sleek.

\*\* NOTE TO SPECIFIER \*\* Delete finial style not required.

* + - * 1. Finial Style: Standard.
        2. Finial Style: Traditional.

\*\* NOTE TO SPECIFIER \*\* Delete one of the two following subparagraphs

* + - * 1. Assembled.
        2. Boxed.
      1. Model LB-HALFC:

\*\* NOTE TO SPECIFIER \*\* Delete frame color not required.

* + - * 1. Frame Color: Stone.
        2. Frame Color: Silver.
        3. Assist Rails: 1/2 Clocking.

\*\* NOTE TO SPECIFIER \*\* Delete headboard and footboard not required

* + - * 1. Headboard and Footboard: Standard.
        2. Headboard and Footboard: Sleek.

\*\* NOTE TO SPECIFIER \*\* Delete finial style not required.

* + - * 1. Finial Style: Standard.
        2. Finial Style: Traditional.

\*\* NOTE TO SPECIFIER \*\* Delete one of the two following subparagraphs

* + - * 1. Assembled.
        2. Boxed.
      1. Standard Features:
         1. Adjustable Length: 80 to 84 in (2032 to 2134 mm).
         2. SlideGuard frame technology.
         3. Auto contour.
         4. Mattress retainer.
         5. 110 voltage.
         6. Vascular lift.
         7. Battery backup.
         8. Pendant control.
         9. Safe Working Load: 450 lbs (204 kg).
         10. Bed Weight: 300 lbs (136 kg).

\*\* NOTE TO SPECIFIER \*\* Delete caster option not required.

* + - * 1. Casters: 3 in (75 mm) diameter.
        2. Casters: None.

\*\* NOTE TO SPECIFIER \*\* Delete assist bar if not required.

* + - * 1. Assist Bar.

\*\* NOTE TO SPECIFIER \*\* Delete finishes not required.

* + - * 1. Finish: V01 Red Oak.
        2. Finish: V02 Solara Oak.
        3. Finish: V03 Newport Apple.
        4. Finish: V05 Clove.
        5. Finish: V06 Acajou Mahogany.
        6. Finish: V07 American Natural.
        7. Finish: V08 Hayward Cherry.
        8. Finish: V31 Cognac Cherry.
      1. Sleep Deck Dimensions:
         1. Length:

Standard Sleep Deck: 93-1/4 in (2369 mm).

Extended Sleep Deck: 97-1/4 in (2471 mm).

Width with Side Rails: 40 in (1016 mm).

* + - 1. Sleep Deck Height Range: 8-1/2 to 26-1/2 in (216 to 673 mm).
      2. Bed Clearance Under Frame: 5-1/2 in (140 mm).
      3. Bed Articulation:
         1. Head Section Inclination: 60 degrees.
         2. Knee Section Inclination: 20 degrees.

\*\* NOTE TO SPECIFIER \*\* Delete paragraph if not required.

* + 1. Basis of Design: Hillrom Smart Beds and Surfaces; Resident Long Term Care Beds.

\*\* NOTE TO SPECIFIER \*\* Delete models not required.

* + - 1. Model LTC-320:
         1. Half head rails with controls.
         2. Half foot rails, no controls.

\*\* NOTE TO SPECIFIER \*\* Delete sleep decks not required.

* + - * 1. Sleep Deck: 76 in (1930 mm) Hard pan.
        2. Sleep Deck: 76 in (1930 mm) Spring fabric.
        3. Sleep Deck: 80 in (2032 mm) Hard pan.
        4. Sleep Deck: 80 in (2032 mm) Spring fabric.
        5. Floor Brakes: Quantity of 4.
      1. Model LTC-310:
         1. Half head rails with controls.

\*\* NOTE TO SPECIFIER \*\* Delete sleep decks not required.

* + - * 1. Sleep Deck: 76 in (1930 mm) Hard pan.
        2. Sleep Deck: 76 in (1930 mm) Spring fabric.
        3. Sleep Deck: 80 in (2032 mm) Hard pan.
        4. Sleep Deck: 80 in (2032 mm) Spring fabric.
        5. Assist bar: Foot end.
        6. Floor Brakes: Quantity of 4.
      1. Model LTC-300:
         1. Half head rails with controls.

\*\* NOTE TO SPECIFIER \*\* Delete sleep decks not required.

* + - * 1. Sleep Deck: 76 in (1930 mm) Hard pan.
        2. Sleep Deck: 76 in (1930 mm) Spring fabric.
        3. Sleep Deck: 80 in (2032 mm) Hard pan.
        4. Sleep Deck: 80 in (2032 mm) Spring fabric.
        5. Assist Bar: Foot end.
        6. Floor Brakes: Quantity of 4.
      1. Model LTC-250:
         1. Half head rails no controls.

\*\* NOTE TO SPECIFIER \*\* Delete sleep decks not required.

* + - * 1. Sleep Deck: 76 in (1930 mm) Hard pan.
        2. Sleep Deck: 76 in (1930 mm) Spring fabric.
        3. Sleep Deck: 80 in (2032 mm) Hard pan.
        4. Sleep Deck: 80 in (2032 mm) Spring fabric.
        5. Pendant.
        6. Assist bar: Foot end.
        7. Floor Brakes: Quantity of 4.
      1. Model LTC-205:

\*\* NOTE TO SPECIFIER \*\* Delete sleep decks not required.

* + - * 1. Sleep Deck: 76 in (1930 mm) Hard pan.
        2. Sleep Deck: 76 in (1930 mm) Spring fabric.
        3. Sleep Deck: 80 in (2032 mm) Hard pan.
        4. Sleep Deck: 80 in (2032 mm) Spring fabric.
        5. Pendant.
        6. Floor Brakes: Quantity of 4.
        7. No head or foot rails.
      1. Model LTC-105:

\*\* NOTE TO SPECIFIER \*\* Delete sleep decks not required.

* + - * 1. Sleep Deck: 76 in (1930 mm) Hard pan.
        2. Sleep Deck: 76 in (1930 mm) Spring fabric.
        3. Sleep Deck: 80 in (2032 mm) Hard pan.
        4. Sleep Deck: 80 in (2032 mm) Spring fabric.
        5. Pendant.
        6. Floor Brakes: Quantity of 2.
        7. No head or foot rails.
      1. Model LTC-100:
         1. Half head rails no controls.

\*\* NOTE TO SPECIFIER \*\* Delete sleep decks not required.

* + - * 1. Sleep Deck: 76 in (1930 mm) Hard pan.
        2. Sleep Deck: 76 in (1930 mm) Spring fabric.
        3. Sleep Deck: 80 in (2032 mm) Hard pan.
        4. Sleep Deck: 80 in (2032 mm) Spring fabric.
        5. Pendant.
        6. Floor Brakes: Quantity of 2.
      1. Standard Features:
         1. Control lockout panel.
         2. Auto contour.
         3. Voltage: 110 Vac.
         4. Casters: 3 in (75 mm) diameter.
         5. Wall bumpers.
         6. Mattress retainer.
         7. Paint Color: Taupe.
         8. Safe Working Load: 400 lbs (181 kg).
         9. Bed Weight: 275 lbs (125 kg).
         10. Headboards and Footboards:

\*\* NOTE TO SPECIFIER \*\* Delete styles not required.

Century Style:

\*\* NOTE TO SPECIFIER \*\* Delete colors not required.

Almond leather.

Antique white.

Dark cherry.

Honey maple.

Light neutral.

Light oak.

Medium oak.

Natural legacy.

Natural maple.

Shaker cherry.

Wild cherry.

White nebula.

Hearthside Style:

\*\* NOTE TO SPECIFIER \*\* Delete colors not required.

Almond leather.

Antique white.

Dark cherry.

Honey maple.

Light neutral.

Light oak.

Medium oak.

Natural legacy.

Natural maple.

Shaker cherry.

Wild cherry.

White nebula.

Coventry Style:

\*\* NOTE TO SPECIFIER \*\* Delete colors not required.

Acajou mahogany.

Almond leather.

American natural.

Antique white.

Clove.

Cognac cherry.

Hayward cherry.

Light neutral.

Montana walnut.

Natural legacy.

Newport apple.

Red oak.

Solara oak.

White nebula.

Turtlecreek Style:

\*\* NOTE TO SPECIFIER \*\* Delete colors not required.

Acajou mahogany.

Almond leather.

American natural.

Antique white.

Clove.

Cognac cherry.

Hayward cherry.

Light neutral.

Montana walnut.

Natural legacy.

Newport apple.

Red oak.

Solara oak.

White nebula.

Jamesburg Style:

\*\* NOTE TO SPECIFIER \*\* Delete colors not required.

Acajou mahogany.

Almond leather.

American natural.

Antique white.

Clove.

Cognac cherry.

Hayward cherry.

Light neutral.

Montana walnut.

Natural legacy.

Newport apple.

Red oak.

Solara oak.

White nebula.

* + - 1. Optional Features:

\*\* NOTE TO SPECIFIER \*\* Delete options not required.

* + - * 1. Pendant control.
        2. Assist bar.
        3. Foot rails.
        4. NP100 surface.
      1. Overall Frame Dimensions:
         1. Width (side rails up): 42 in (1067 mm).
         2. Length:

Bumpers in High Position: 95-1/4 in (2419 mm).

Bumpers in Low Position: 91-1/2 in (2324 mm).

* + - 1. Sleep Deck Dimensions:
         1. Width: 36 in (914 mm).
         2. Length: 80 in (2032 mm).
      2. Sleep Deck Height Range: 13-7/8 to 30 in (352 to 762 mm) not including mattress.
      3. Bed Clearance Under Frame: 5 in (127 mm).
      4. Bed Articulations:
         1. Head Elevation: 67 degrees.
         2. Knee Elevations: 27 degrees.

\*\* NOTE TO SPECIFIER \*\* Delete article if not required.

* 1. STRETCHERS

\*\* NOTE TO SPECIFIER \*\* Delete paragraph not required.

* + 1. Basis of Design: Hillrom Safe Patient Handling and Mobility; Model P8000 Procedural Stretcher (PS).
       1. Standard Compliance:
          1. CSA C22.2 No. 601.1.
          2. EN 60601-1.
          3. ISO 10993-1.
          4. ISO 10993-5.
          5. ISO 10993-10.
          6. ISO 14971.
       2. Standard Features:
          1. Capacity: 700 lbs (317.5 kg) safe working load.
          2. Height: 20.7 in (526 mm).
          3. Trendelenburg and Reverse Trendelenburg: 18 degrees.
          4. Full perimeter bumpering.
          5. Standard IV Pole Sockets: 3/4 in (19 mm).
          6. Hydraulic Control Pedals: Both sides.
          7. Transport Strap Locations: Six (trauma - 4).
          8. Storage area.
          9. Pneumatic head section.
          10. Casters: 8 in (203 mm) diameter.
          11. Bed Weight: 265 lbs (120.2 kg).
          12. Power washable.
       3. Frame Options:

\*\* NOTE TO SPECIFIER \*\* Delete frame options not required.

* + - * 1. High impact shroud with oxygen tank storage.
        2. Dual display integrated scale.
        3. Integrated handbrake system.
      1. Deck Features:

\*\* NOTE TO SPECIFIER \*\* Delete deck width not required.

* + - * 1. Deck Width: 26 in (660 mm). Standard.
        2. Deck Width: 30 in (762 mm).

\*\* NOTE TO SPECIFIER \*\* Delete sleep surface not required. Hard pan non-radiolucent is standard.

* + - * 1. Sleep Surface: Hard pan non-radiolucent.
        2. Sleep Surface: Radiolucent.
        3. Sleep Surface: Radiolucent with upright cassette holder.

\*\* NOTE TO SPECIFIER \*\* Delete push handles not required.

* + - * 1. Push Handles: At head.
        2. Push Handles: At head with integrated O2 tank holder with utility tray and blue handles.
        3. Push Handles: At foot.
        4. Push Handles: At head and foot.
        5. Push Handles: At head and foot with integrated O2 tank holder with utility tray and blue handles.

\*\* NOTE TO SPECIFIER \*\* Delete pole locations not required.

* + - * 1. 3 Stage IV Pole Locations: At head right side.
        2. 3 Stage IV Pole Locations: At foot left side.
        3. 3 Stage IV Pole Locations: At head and foot.
      1. Deck Options:

\*\* NOTE TO SPECIFIER \*\* Delete options not required.

* + - * 1. Knee gatch.
        2. Backsaver Fowler feature (comes with Knee gatch).
        3. Auto Contour with Knee gatch and Backsaver Fowler.
        4. Hydraulic Pedals: Both sides and foot.
      1. Side Rail Features:
         1. Siderail End Gap: At head.
         2. One-step TuckAway Siderails for aero-gap transfers.
      2. Side Rail Options:

\*\* NOTE TO SPECIFIER \*\* Delete options not required.

* + - * 1. Siderail end gap at foot.
        2. KX1 Additional Coverage Siderail: Foot end gap only.
      1. Steering Features:
         1. Steering Plus System: Uses 5th wheel as pivot.

\*\* NOTE TO SPECIFIER \*\* Delete brake and steer pedals not required. At head and foot ends is standard.

* + - * 1. Brake and Steer Pedals: At head and foot ends.
        2. Brake and Steer Pedals: At head and foot ends, and both sides. Four sided.
      1. Integrated IV Pole Transfer Handles Options:

\*\* NOTE TO SPECIFIER \*\* Delete options not required.

* + - * 1. Both Ends: Right and left side.
        2. Both Ends: Right side.
        3. Both Ends: Left side.
        4. Foot End: Right and left side.
        5. Foot End: Right side.
        6. Foot End: Left side.
        7. Head End: Right and left side.
        8. Head End: Right side.
        9. Head End: Left side.
      1. Bumper Color Options:

\*\* NOTE TO SPECIFIER \*\* Delete bumper color options not required.

* + - * 1. Light neutral.
        2. Blue.
        3. Teal.
        4. Purple.
        5. Red.
      1. Mattress: Mattresses comply with north American federal regulations.

\*\* NOTE TO SPECIFIER \*\* Delete mattresses and mattress dimensions not required.

* + - * 1. Comfortline Foam Mattress with Fire Barrier:

Dimensions (T x W x L): 3 x 26 x 75 in (77 x 660 x 1905 mm). Standard.

Dimensions (T x W x L): 3 x 30 x 75 in (77 x 762 x 1905 mm).

* + - * 1. High Density Foam Mattress with Fire Barrier:

Dimensions (T x W x L): 4 x 26 x 75 in (102 x 660 x 1905 mm).

Dimensions (T x W x L): 4 x 30 x 75 in (102 x 762 x 1905 mm).

* + - * 1. Comfortline Multi-Density Foam Mattress with Fire Barrier:

Dimensions (T x W x L): 5 x 26 x 75 in (127 x 660 x 1905 mm).

Dimensions (T x W x L): 5 x 30 x 75 in (127 x 762 x 1905 mm).

* + - * 1. AccuMax Quantum Elastimax Ticking:

Dimensions (T x W x L): 5 x 26 x 75 in (127 x 660 x 1905 mm).

Dimensions (T x W x L): 5 x 30 x 75 (127 x 762 x 1905 mm).

* + - * 1. AccuMax Quantum Hydroplex Ticking:

Dimensions (T x W x L): 5 x 26 x 75 in (127 x 660 x 1905 mm).

Dimensions (T x W x L): 5 x 30 x 75 (127 x 762 x 1905 mm).

* + - 1. Decal Options:

\*\* NOTE TO SPECIFIER \*\* Delete decals options not required.

* + - * 1. Emergency.
        2. Endoscopy.
        3. Cath Lab.
        4. GI Lab.
        5. PACU.
        6. Transport.
        7. Maternity.
        8. Radiology.
        9. Surgery.
        10. Ambulatory Surgery.
      1. Dimensional Parameters:
         1. Overall Length: 83 in (2108 mm).
         2. Overall Width (Siderails Up): 32 in (813 mm).
         3. Overall Width (Siderails Up): 36 in (914 mm).
         4. Overall Width (Siderails Stored): 26.5 in (673 mm).
         5. Overall Width (Siderails Stored): 30.25 in (768 mm).
         6. Siderail Length: 47 in (1194 mm).
         7. Siderail Height Above Sleep Deck: 14.5 in (368 mm).
         8. Maximum Head Elevation: 90 degrees.
         9. Floor to Base Clearance: 3.5 in (89 mm) without power transport module.

\*\* NOTE TO SPECIFIER \*\* Delete paragraph not required.

* + 1. Basis of Design: Hillrom Safe Patient Handling and Mobility; Model P8000 Intellidrive Stretcher (PSS).
       1. Standard Compliance:
          1. CSA C22.2 No. 601.1.
          2. EN 60601-1.
          3. IEC 60601-1-2.
          4. IEC 60601-1-4.
          5. IEC 60601-2-38.
          6. ISO 10993-1.
          7. ISO 10993-5.
          8. ISO 10993-10.
          9. ISO 14971.
       2. Standard Features:
          1. Capacity: 700 lbs (317.5 kg) safe working load.
          2. Height: 20.7 in (526 mm).
          3. Trendelenburg and Reverse Trendelenburg: 12 degrees.
          4. Full perimeter bumpering.
          5. Standard IV Pole Sockets: 3/4 in (19 mm).
          6. Hydraulic Control Pedals: Both sides.
          7. Transport Strap Locations: Six (trauma - 4).
          8. Storage area.
          9. Pneumatic head section.
          10. Casters: 8 in (203 mm) diameter.
          11. Bed Weight: 400 lbs (200 kg).
          12. Power washable.
          13. Electric: 120 Vac.
       3. Frame Features:
          1. IntelliDrive powered transport mechanism.
          2. High impact shroud with oxygen tank storage.

\*\* NOTE TO SPECIFIER \*\* Delete frame option if not required.

* + - * 1. Deck Width: 26 in (660 mm). Standard.
        2. Deck Width: 30 in (762 mm).

\*\* NOTE TO SPECIFIER \*\* Delete sleep surface not required. Hard pan non-radiolucent is standard.

* + - * 1. Sleep Surface: Hard pan non-radiolucent
        2. Sleep Surface: Radiolucent.

\*\* NOTE TO SPECIFIER \*\* Delete push handles not required

* + - * 1. Push Handles: At head.
        2. Push Handles: At head with integrated O2 tank holder with utility tray and blue handles.

\*\* NOTE TO SPECIFIER \*\* Delete pole locations not required

* + - * 1. 3 Stage IV Pole Locations: At head right side.
        2. 3 Stage IV Pole Locations: At head and foot.
      1. Deck Options:

\*\* NOTE TO SPECIFIER \*\* Delete deck options not required.

* + - * 1. Knee gatch.
        2. Backsaver Fowler feature (comes with Knee gatch).
        3. Auto Contour with Knee gatch and Backsaver Fowler feature.
        4. Hydraulic Pedals on Both Sides and Foot.
      1. Side Rail Features:
         1. Siderail End Gap: At head.
         2. One-step TuckAway Siderails for aero-gap transfers.
      2. Steering Feature: Brake and steer pedals at head and foot ends. Four-sided.
      3. Integrated IV Pole Transfer Handles Options:

\*\* NOTE TO SPECIFIER \*\* Delete options not required.

* + - * 1. Head End - Right and Left Side.
        2. Head End - Right Side.
        3. Head End - Left Side.
      1. Bumper Color: Green Bumper and Dark Blue Bellows.
      2. Mattress: Mattresses comply with North American federal regulations.

\*\* NOTE TO SPECIFIER \*\* Delete mattresses not required.

* + - * 1. Comfortline Foam Mattress with Fire Barrier:

Dimensions (T x W x L): 3 x 26 x 75 in (77 x 660 x 1905 mm). Standard.

Dimensions (T x W x L): 3 x 30 x 75 in (77 x 762 x 1905 mm).

* + - * 1. High Density Foam Mattress with Fire Barrier:

Dimensions (T x W x L): 4 x 26 x 75 in (102 x 660 x 1905 mm).

Dimensions (T x W x L): 4 x 30 x 75 in (102 x 762 x 1905 mm).

* + - * 1. Comfortline Multi-Density Foam Mattress with Fire Barrier:

Dimensions (T x W x L): 5 x 26 x 75 in (127 x 660 x 1905 mm).

Dimensions (T x W x L): 5 x 30 x 75 in (127 x 762 x 1905 mm).

* + - * 1. AccuMax Quantum Elastimax Ticking:

Dimensions (T x W x L): 5 x 26 x 75 in (127 x 660 x 1905 mm).

Dimensions (T x W x L): 5 x 30 x 75 (127 x 762 x 1905 mm).

* + - * 1. AccuMax Quantum Hydroplex Ticking:

Dimensions (T x W x L): 5 x 26 x 75 in (127 x 660 x 1905 mm).

Dimensions (T x W x L): 5 x 30 x 75 (127 x 762 x 1905 mm).

* + - 1. Decal Options:

\*\* NOTE TO SPECIFIER \*\* Delete decals options not required.

* + - * 1. Emergency.
        2. Endoscopy.
        3. Cath Lab.
        4. GI Lab.
        5. PACU.
        6. Transport.
        7. Maternity.
        8. Radiology.
        9. Surgery.
        10. Ambulatory Surgery.
      1. Dimensional Parameters:
         1. Overall Length: 84.5 in (2146 mm).
         2. Overall Width (Siderails Up): 32 in (813 mm) or 36 in (914 mm).
         3. Overall Width (Siderails Stored): 29.5 in (743 mm) or 33.25 in (844 mm).
         4. Siderail Length: 47 in (1194 mm).
         5. Siderail Height Above Sleep Deck: 14.5 in (368 mm).
         6. Maximum Head Elevation: 90 degrees.
         7. Floor to Base Clearance: 3.5 in (89 mm) without power transport module.
         8. Below Power Drive Module: 1.5 in (38 mm).

\*\* NOTE TO SPECIFIER \*\* Delete paragraph not required.

* + 1. Basis of Design: Hillrom Safe Patient Handling and Mobility; Model P8005 Transport Stretcher (TN).
       1. Standard Compliance:
          1. CSA C22.2 No. 601.1.
          2. EN 60601-1.
          3. ISO 10993-1.
          4. ISO 10993-5.
          5. ISO 10993-10.
          6. ISO 14971.
       2. Standard Features:
          1. Capacity: 700 lbs (317.5 kg) safe working load.
          2. Trendelenburg and Reverse Trendelenburg: 18 degrees.
          3. Full perimeter bumpering.
          4. Standard IV Pole Sockets: 3/4 in (19 mm).
          5. Hydraulic Control Pedals: Both sides.
          6. Transport Strap Locations: Six (trauma - 4).
          7. Storage area.
          8. Pneumatic head section.
          9. Casters: 8 in (203 mm) diameter.
          10. Bed Weight: 265 lbs (120.2 kg).
       3. Deck Features:
          1. Deck Width: 26 in (660 mm).
          2. Sleep Surface: Hard pan non-radiolucent.
          3. Push handles at foot.
          4. 3 Stage IV pole at foot, left side.
       4. Deck Options:

\*\* NOTE TO SPECIFIER \*\* Delete options not required.

* + - * 1. Knee gatch.
        2. Backsaver Fowler feature (comes with knee gatch).
        3. Auto Contour with knee gatch and Backsaver Fowler.
        4. Hydraulic Pedals: Both sides and foot.
      1. Side Rail Features:
         1. Siderail End Gap: At foot.
         2. One-step TuckAway Siderails for aero-gap transfers.
      2. Steering Features:
         1. Steering Plus System (uses 5th wheel as pivot).
         2. Brake and Steer Pedals: Head and foot ends.
         3. Brake and Steer Pedals: All four sides.
      3. Integrated IV Pole Transfer Handles Options:

\*\* NOTE TO SPECIFIER \*\* Delete options not required.

* + - * 1. Foot End: Right and left side.
        2. Head End: Right and left side.
        3. Foot End: Right side.
        4. Foot End: Left side.
      1. Bumper Color Options:

\*\* NOTE TO SPECIFIER \*\* Delete bumper color options not required.

* + - * 1. Light neutral.
        2. Blue.
        3. Teal.
        4. Purple.
        5. Red.
      1. Mattress: Mattresses comply with north american federal regulations.

\*\* NOTE TO SPECIFIER \*\* Delete mattresses and mattress dimensions not required.

* + - * 1. Comfortline Foam Mattress with Fire Barrier:

Dimensions (T x W x L): 3 x 26 x 75 in (77 x 660 x 1905 mm). Standard.

Dimensions (T x W x L): 3 x 30 x 75 in (77 x 762 x 1905 mm).

* + - * 1. High Density Foam Mattress with Fire Barrier:

Dimensions (T x W x L): 4 x 26 x 75 in (102 x 660 x 1905 mm).

Dimensions (T x W x L): 4 x 30 x 75 in (102 x 762 x 1905 mm).

* + - * 1. Comfortline Multi-Density Foam Mattress with Fire Barrier:

Dimensions (T x W x L): 5 x 26 x 75 in (127 x 660 x 1905 mm).

Dimensions (T x W x L): 5 x 30 x 75 in (127 x 762 x 1905 mm).

* + - * 1. AccuMax Quantum Elastimax Ticking:

Dimensions (T x W x L): 5 x 26 x 75 in (127 x 660 x 1905 mm).

Dimensions (T x W x L): 5 x 30 x 75 (127 x 762 x 1905 mm).

* + - * 1. AccuMax Quantum Hydroplex Ticking:

Dimensions (T x W x L): 5 x 26 x 75 in (127 x 660 x 1905 mm).

Dimensions (T x W x L): 5 x 30 x 75 (127 x 762 x 1905 mm).

* + - 1. Decal Options:

\*\* NOTE TO SPECIFIER \*\* Delete decals not required.

* + - * 1. Emergency.
        2. Endoscopy.
        3. Cath Lab.
        4. GI Lab.
        5. PACU.
        6. Transport.
        7. Maternity.
        8. Radiology.
        9. Surgery.
        10. Ambulatory Surgery.

\*\* NOTE TO SPECIFIER \*\* Delete paragraph not required.

* + 1. Basis of Design: Hillrom Safe Patient Handling and Mobility; Model P8010 Surgical Stretcher (SS).
       1. Standard Compliance:
          1. CSA C22.2 No. 601.1.
          2. EN 60601-1.
          3. IEC 60601-2-38.
          4. IEC 60601-2-46.
          5. ISO 10993-1.
          6. ISO 10993-5.
          7. ISO 10993-10.
          8. ISO 14971.
       2. Standard Features:
          1. Capacity: 700 lbs (317.5 kg) maximum patient weight.
          2. Trendelenburg and Reverse Trendelenburg: 18 degrees.
          3. Full perimeter bumpering.
          4. Standard IV Pole Sockets: 3/4 in (19 mm).
          5. Hydraulic Control Pedals: Both sides.
          6. Transport Strap Locations: Six (trauma - 4).
          7. Large storage area.
          8. Pneumatic head section.
          9. Casters: 8 in (203 mm) diameter.
          10. Bed Weight: 285 lbs (129.3 kg)
       3. Deck Features:
          1. Deck Width: 26 in (660 mm). Standard.
          2. Deck Width: 30 in (762 mm).
          3. Hard Pan Non-Radiolucent Sleep Surface.
       4. Deck Options:
          1. Radiolucent Head:

\*\* NOTE TO SPECIFIER \*\* Delete if not required.

Upright cassette holder.

\*\* NOTE TO SPECIFIER \*\* Always retain one option below. Delete options not required.

* + - * 1. Push Handles: At head.
        2. Push Handles: At foot.
        3. Push Handles: At head and foot.
        4. Integrated O2 tank holder, utility tray with blue handles.
        5. 3 Stage IV Pole: At head right side.
        6. 3 Stage IV Pole: At foot left side.
        7. 3 Stage IV Pole: At head and foot.
      1. Side Rail Features:
         1. Siderail End Gap: At head.

\*\* NOTE TO SPECIFIER \*\* Delete if not required.

* + - * 1. Fold-down side rails.

\*\* NOTE TO SPECIFIER \*\* Delete options not required,

* + - 1. Steering Features:

\*\* NOTE TO SPECIFIER \*\* Delete pivoting feature not required.

* + - * 1. Pivoting: Locking steering caster as pivot.
        2. Pivoting: Steering Plus System using 5th wheel as pivot.

\*\* NOTE TO SPECIFIER \*\* Delete brake and steering pedals not required.

* + - * 1. Brake and Steer Pedals: At head and foot ends.
        2. Brake and Steer Pedals: All four sides.
      1. Integrated IV Pole Transfer Handles Options:

\*\* NOTE TO SPECIFIER \*\* Delete head end not required.

* + - * 1. Head End: Right side.
        2. Head End: Left side.
      1. Bumper Color Options:

\*\* NOTE TO SPECIFIER \*\* Delete colors not required.

* + - * 1. Blue.
        2. Teal.
        3. Purple.
        4. Red.
        5. Taupe.
      1. Mattress Options (Mattresses comply with North American Federal Regulations):
         1. High Density Foam Mattress with Fire Barrier:

\*\* NOTE TO SPECIFIER \*\* Delete surgical surface not required.

Surgical surface-2 pc flat/concave head cushion set.

Dimensions (T x W x L): 4 x 26 x 78 in (102 x 660 x 1981 mm).

Surgical surface with flat head cushion set.

Dimensions (T x W x L): 4 x 26 x 78 in (102 x 660 x 1981 mm).

* + - 1. Decal Options:

\*\* NOTE TO SPECIFIER \*\* Delete decals not required.

* + - * 1. Emergency.
        2. Endoscopy.
        3. Cath Lab.
        4. GI Lab.
        5. PACU.
        6. Transport.
        7. Maternity.
        8. Radiology.
        9. Surgery.
        10. Ambulatory Surgery.
      1. Technical Specifications:
         1. Overall Length: 92 in (2337 mm).
         2. Overall Width (Siderails Up): 32 in (813 mm).
         3. Overall Width (Siderails Stored): 26 in (660 mm).
         4. Siderail Length: 47 in (1194 mm).
         5. Siderail Height Above Sleep Deck: 14.5 in (368 mm).
         6. Maximum Head Elevation: 70 degrees.
         7. Floor to Base Clearance: 3.5 in (89 mm) without power transport module.

\*\* NOTE TO SPECIFIER \*\* Delete paragraph not required.

* + 1. Basis of Design: Hillrom Safe Patient Handling and Mobility; Model P8020 Electric Stretcher (ES).
       1. Standard Compliance:
          1. CSA C22.2 No. 601.1.
          2. EN 60601-1.
          3. IEC 60601-1-2.
          4. IEC 60601-2-38.
          5. ISO 10993-1.
          6. ISO 10993-5.
          7. ISO 10993-10.
          8. ISO 14971.
       2. Standard Features:
          1. Capacity: 700 lbs (317.5 kg) safe working load.
          2. Trendelenburg and Reverse Trendelenburg.
          3. Full perimeter bumpering.
          4. Standard IV Pole Sockets: 3/4 in (19 mm).
          5. Hydraulic Control Pedals: Both sides.
          6. Transport Strap Locations: Six (trauma - 4).
          7. Storage area.
          8. Pneumatic head section.
          9. Casters: 8 in (203 mm) diameter.
          10. Bed Weight: 355 lbs (161.0 kg).
          11. Electric: 120 Vac.
       3. Frame Options:

\*\* NOTE TO SPECIFIER \*\* Delete frame options not required.

* + - * 1. Dual display integrated scale.
        2. Integrated handbrake system.
      1. Deck Features:
         1. Deck Width: 30 in (762 mm).
         2. Hard Pan Non-Radiolucent Sleep Surface.
         3. Knee gatch.

\*\* NOTE TO SPECIFIER \*\* Delete push handles not required.

* + - * 1. Push Handles: At head.
        2. Push Handles: At foot.
        3. Push Handles: At head and foot.

\*\* NOTE TO SPECIFIER \*\* Delete pole locations not required.

* + - * 1. 3 Stage IV Pole Locations: At head right side.
        2. 3 Stage IV Pole Locations: At foot left side.
        3. 3 Stage IV Pole Locations: At head and foot.
      1. Siderail Features:
         1. Siderail End Gap: At head.
         2. One-step TuckAway siderails for aero-gap transfers.
      2. Steering Features:
         1. Steering Plus System: Uses 5th wheel as pivot.

\*\* NOTE TO SPECIFIER \*\* Delete brake and steer pedals not required. At head and foot ends is standard.

* + - * 1. Brake and Steer Pedals: At head and foot ends.
        2. Brake and Steer Pedals: At head and foot ends, and both sides. Four sided.
      1. Integrated IV Pole Transfer Handles Options:

\*\* NOTE TO SPECIFIER \*\* Delete options not required.

* + - * 1. Both Ends: Right and left side.
        2. Both Ends: Right side.
        3. Both Ends: Left side.
        4. Foot End: Right and left side.
        5. Foot End: Right side.
        6. Foot End: Left side.
        7. Head End: Right and left side.
        8. Head End: Right side.
        9. Head End: Left side.
      1. Bumper Color Options:
         1. Light neutral.
         2. Blue.
         3. Teal.
         4. Purple.
         5. Red.
      2. Mattress Options (Mattresses comply with North American Federal Regulations):

\*\* NOTE TO SPECIFIER \*\* Delete mattresses not required.

* + - * 1. High Density Foam Mattress with Fire Barrier:

Dimensions (T x W x L): 4 x 26 x 75 in (102 x 660 x 1905 mm).

Dimensions (T x W x L): 4 x 30 x 75 in (102 x 762 x 1905 mm).

* + - * 1. Comfortline Multi-Density Foam Mattress with Fire Barrier:

Dimensions (T x W x L): 5 x 30 x 75 in (127 x 762 x 1905 mm).

* + - * 1. AccuMax Quantum Elastimax Ticking:'

Dimensions (T x W x L): 5 x 30 x 75 (127 x 762 x 1905 mm).

* + - * 1. AccuMax Quantum Hydroplex Ticking:

Dimensions (T x W x L): 5 x 30 x 75 (127 x 762 x 1905 mm).

* + - 1. Decal Options:

\*\* NOTE TO SPECIFIER \*\* Delete decals not required.

* + - * 1. Emergency.
        2. Endoscopy.
        3. Cath Lab.
        4. GI Lab.
        5. PACU.
        6. Transport.
        7. Maternity.
        8. Radiology.
        9. Surgery.
        10. Ambulatory Surgery.
      1. Dimensional Parameters:
         1. Overall Length: 83 in (2108 mm).
         2. Overall Width (Siderails Up): 36 in (914 mm).
         3. Overall Width (Siderails Stored): 34 in (864 mm).
         4. Siderail Length: 47 in (1194 mm).
         5. Siderail Height Above Sleep Deck: 14.5 in (368 mm).
         6. Maximum Head Elevation: 90 degrees.

\*\* NOTE TO SPECIFIER \*\* Delete paragraph not required.

* + 1. Basis of Design: Hillrom Safe Patient Handling and Mobility; Model P8040 Trauma Stretcher (TS).
       1. Standard Compliance:
          1. CSA C22.2 No. 601.1.
          2. EN 60601-1.
          3. ISO 10993-1.
          4. ISO 10993-5.
          5. ISO 10993-10.
          6. ISO 14971.
       2. Standard Features:
          1. Capacity: 700 lbs (317.5 kg) safe working load.
          2. Trendelenburg and Reverse Trendelenburg: 18 degrees.
          3. Full perimeter bumpering.
          4. Standard IV Pole Sockets: 3/4 in (19 mm).
          5. Hydraulic Control Pedals: Both sides.
          6. Transport Strap Locations: Six (trauma - 4).
          7. Storage area.
          8. Pneumatic head section.
          9. Casters: 8 in (203 mm) diameter.
          10. Bed Weight: 290 lbs (131.5 kg).
          11. Power washable.
       3. Frame Options:
          1. Dual display integrated scale.
          2. Integrated handbrake system.
       4. Deck Features
          1. Deck Width: 26 in (660 mm). Standard.
          2. Deck Width: 30 in (762 mm).

\*\* NOTE TO SPECIFIER \*\* Delete radiolucent surface not required.

* + - * 1. Radiolucent surface.
        2. Radiolucent surface with upright cassette holder.

\*\* NOTE TO SPECIFIER \*\* Delete push handles not required

* + - * 1. Push Handles: At head.
        2. Push Handles: At foot.
        3. Push Handles: At head and foot.

\*\* NOTE TO SPECIFIER \*\* Delete pole locations not required

* + - * 1. 3 Stage IV Pole Locations: At head right side.
        2. 3 Stage IV Pole Locations: At foot left side.
        3. 3 Stage IV Pole Locations: At head and foot.

\*\* NOTE TO SPECIFIER \*\* Optional. Delete if not required.

* + - * 1. Hydraulic pedals on both sides and foot.
      1. Siderail Features:
         1. One-Step TuckAway siderails for zero-gap transfers.

\*\* NOTE TO SPECIFIER \*\* Delete siderails not required. At head is standard.

* + - * 1. Siderail End Gap: At head.
        2. Siderail End Gap: At foot.
        3. Siderail End Gap: At foot with KX1 additional c overage siderail.
      1. Steering Features:
         1. Steering Plus System: Uses 5th wheel as pivot.

\*\* NOTE TO SPECIFIER \*\* Delete brake and steer pedals not required. At head and foot ends is standard.

* + - * 1. Brake and Steer Pedals: At head and foot ends.
        2. Brake and Steer Pedals: At head and foot ends, and both sides. Four sided.
      1. Integrated IV Pole Transfer Handles Options:

\*\* NOTE TO SPECIFIER \*\* Delete options not required.

* + - * 1. Both Ends: Right and left side.
        2. Both Ends: Right side.
        3. Both Ends: Left side.
        4. Foot End: Right and left side.
        5. Foot End: Right side.
        6. Foot End: Left side.
        7. Head End: Right and left side.
        8. Head End: Right side.
        9. Head End: Left side.
      1. Bumper Color Options

\*\* NOTE TO SPECIFIER \*\* Delete bumper color options not required.

* + - * 1. Light neutral.
        2. Blue.
        3. Teal.
        4. Purple.
        5. Red.
      1. Mattress Options (Mattresses comply with North American Federal Regulations):

\*\* NOTE TO SPECIFIER \*\* Delete mattresses and mattress dimensions not required.

* + - * 1. Comfortline Foam Mattress with Fire Barrier:

Dimensions (T x W x L): 3 x 26 x 75 in (77 x 660 x 1905 mm). Standard.

Dimensions (T x W x L): 3 x 30 x 75 in (77 x 762 x 1905 mm).

* + - * 1. High Density Foam Mattress with Fire Barrier:

Dimensions (T x W x L): 4 x 26 x 75 in (102 x 660 x 1905 mm).

Dimensions (T x W x L): 4 x 30 x 75 in (102 x 762 x 1905 mm).

* + - * 1. Comfortline Multi-Density Foam Mattress with Fire Barrier:

Dimensions (T x W x L): 5 x 26 x 75 in (127 x 660 x 1905 mm).

Dimensions (T x W x L): 5 x 30 x 75 in (127 x 762 x 1905 mm).

* + - * 1. AccuMax Quantum Elastimax Ticking:

Dimensions (T x W x L): 5 x 26 x 75 in (127 x 660 x 1905 mm).

Dimensions (T x W x L): 5 x 30 x 75 (127 x 762 x 1905 mm).

* + - * 1. AccuMax Quantum Hydroplex Ticking:

Dimensions (T x W x L): 5 x 26 x 75 in (127 x 660 x 1905 mm).

Dimensions (T x W x L): 5 x 30 x 75 (127 x 762 x 1905 mm).

* + - 1. Decal Options:
         1. Emergency.
         2. Endoscopy.
         3. Cath Lab.
         4. GI Lab.
         5. PACU.
         6. Transport.
         7. Maternity.
         8. Radiology.
         9. Surgery.
         10. Ambulatory Surgery.
      2. Dimensional Parameters:
         1. Overall Length: 83 in (2108 mm).
         2. Overall Width (Siderails Up): 32 in (813 mm).
         3. Overall Width (Siderails Stored): 26.5 in (673 mm).
         4. Siderail Length: 47 in (1194 mm).
         5. Siderail Height Above Sleep Deck: 11 in (279 mm).
         6. Maximum Head Elevation: 90 degrees.
         7. Floor to Base Clearance: 3.5 in (89 mm) without power transport module.

\*\* NOTE TO SPECIFIER \*\* Delete paragraph not required.

* + 1. Basis of Design: Hillrom Safe Patient Handling and Mobility; Model P8050 OB-GYN (OS).
       1. Standard Compliance:
          1. CSA C22.2 No. 601.1.
          2. EN 60601-1.
          3. IEC 60601-2-38.
          4. ISO 10993-1.
          5. ISO 10993-5.
          6. ISO 10993-10.
          7. ISO 14971.
       2. Standard Features:
          1. Capacity: 700 lbs (317.5 kg) maximum patient eight.
          2. Height: 20.7 in (526 mm).
          3. Trendelenburg and Reverse Trendelenburg: 18 degrees.
          4. Full perimeter bumpering.
          5. Standard IV Pole Sockets: 3/4 in (19 mm).
          6. Hydraulic Control Pedals: Both sides.
          7. Transport Strap Locations: Six (trauma - 4).
          8. Storage area.
          9. Pneumatic head section.
          10. Casters: 8 in (203 mm) diameter.
          11. Bed Weight: 325 lbs (147.4 kg).
       3. Deck Features:
          1. Width: 30 in (762 mm).

\*\* NOTE TO SPECIFIER \*\* Delete radiolucent surface not required.

* + - * 1. Radiolucent surface.
        2. Radiolucent surface with upright cassette holder.
      1. Deck Options:
         1. Push Handles: At head.
         2. 3 Stage IV Pole Locations: At head right side.
      2. Siderail Features:
         1. Siderail End Gap: At head.
         2. One-step TuckAway siderails for aero-gap transfers.
      3. Steering Features:
         1. Steering Plus System: Uses 5th wheel as pivot.

\*\* NOTE TO SPECIFIER \*\* Delete brake and steer pedals not required. At head and foot ends is standard.

* + - * 1. Brake and Steer Pedals: At head and foot ends.
        2. Brake and Steer Pedals: At head and foot ends, and both sides. Four sided.
      1. Integrated IV Pole Transfer Handles Options:

\*\* NOTE TO SPECIFIER \*\* Delete head end not required.

* + - * 1. Head End: Right side.
        2. Head End: Left side.
      1. Bumper Color Options:
         1. Light neutral.
         2. Blue.
         3. Teal.
         4. Purple.
         5. Red.
      2. Mattress: 3 in (76 mm) thick, OB-GYN.
      3. Decals Options:
         1. Emergency.
         2. Endoscopy.
         3. Cath Lab.
         4. GI Lab.
         5. PACU.
         6. Transport.
         7. Maternity.
         8. Radiology.
         9. Surgery.
         10. Ambulatory Surgery.
      4. Technical Specifications:
         1. Overall Length: 83 in (2108 mm).
         2. Overall Width (Siderails Up): 32 in (813 mm).
         3. Overall Width (Siderails Stored): 26.5 in (673 mm).
         4. Siderail Length: 47 in (1194 mm).
         5. Siderail Height Above Sleep Deck: 11 in (279 mm).
         6. Maximum Head Elevation: 90 degrees.
         7. Floor to Base Clearance: 3.5 in (89 mm) without power transport module.

1. EXECUTION
   1. EXAMINATION
      1. Verification of Conditions: Examine areas and conditions under which products listed in this specification are to be used to assure products will perform as specified and identify conditions that may be detrimental to proper performance.
      2. Do not proceed until unsatisfactory conditions have been corrected.
   2. INSTALLATION
      1. General: Install per manufacturer's written instructions and in proper relationship with adjacent construction.
      2. Testing: Test per manufacturer's written instructions. Adjust until satisfactory results are obtained. If satisfactory results cannot be obtained, replace unit at no additional cost to the Owner.
   3. CLEANING AND PROTECTION
      1. Protect from damage during construction operations. Promptly repair any damaged surfaces. Remove and replace work which cannot be satisfactorily repaired.
      2. Clean products, prior to Substantial Completion, using materials recommended by the manufacturer.

END OF SECTION